OhmPi monitoring installation checklist
Date / e.g YYYY-MM-DD
Installing team / installer name(s)
Location / e.g. Rocherfort (BE), 54.2223, 3.4423
OhmPi name / register your ohmpi on https://ohmpi.org
OhmPi config / e.g. mb2024+4*mux2023+dph5005

Prior to field installation (in the lab)

Measurement board
  - Continuity check (SC checks) (power off)
  - Voltage check (power on)
  - Check shunt resistor, board version and values in config.py
  - Resistance check on reference resistor board with quad [0,0,0,0]
- Multiplexer board
  - All role cables (A, B, M, N) wired to right board/connector and MUX boards addresses match those in config.py
  - Continuity test (OhmPi.test_mux())
  - Resistance check on reference resistor board with reference_sequence file
In the field (system not powered)
- Immediate surrounding free of power sources that could affect the system

  e.g. generator, high voltage line, etc.
- Cables firmly connected to screw terminals (gently pull wires to check connections)
- Components firmly inserted in their socket (e.g. Traco power, MCP23008, Current 7 click etc.)
- Metallic enclosure wired to ground
- System protected from rainfall/sun and humidity
  e.g. sun shade, silicagel, water tight spray, etc.
- Charged batteries
- Decoupled solar charging system from battery when measuring
Power the system
Software tests (OhmPi.test())

Resistance check on reference resistor board with reference_sequence file
Connect your electrode arrays
- Contact resistance check

- Sequence run
- Current does not reach the limit of 4800 / 50 /shunt_value mA
  e.g. for a shunt of 2 Ohm, this means a limit of 48 mA
- Full-waveform check for noise or strange decay (OhmPi.plot_last_fw())

Installation completed